## Experiment Overview
2025-11-05_095412_DEN078_RNaseP-16S.eds
| Experiment Aims |  |  |  | Reagents Used
1 | DEN078 Rnase P-16s qPCR |  |  |  | Component |  | Supplier | Cat. No. | Lot No.
2 |  |  |  |  | TaqMan® Fast Advanced Master Mix (5x5mL) |  | Fisher Scientific | 13438456 | 4444554
3 |  |  |  |  | RNase P Primer-Probe (FAM) |  | Thermofisher | 4316831 | 3245729
4 |  |  |  |  | Human Genomic DNA |  | Promega | G3041 | 0000636473
5 |  |  |  |  | Human DNA Male |  | Thermofisher | 360486 | 3106717
6 |  |  |  |  | TaqMan® Gene Expression Assay (750rxns)* |  | Fisher Scientific | 11957021 | P250904-011-F09
7 |  |  |  |  | Microbial Community DNA Standard (2000ng) |  | ZymoBIOMICS | D6306 | 24027229
[5 empty] | Nuclease free water or suitable alternative |  | QIAGEN | 1039498 | N/A
Batch ID: | DEN078 |  |  |  | TE (1X) buffer (pH8.0) or suitable alternative |  | Fisher Scientific | 10224683 | 250226
Operator: | KP
Date: | 2025-12-01
Equipment Used
| Asset # | Item | Make | Model |  | Serial number
| 60 | Laminar Flow Cabinet | StarLab | GuardOne |  | 9206IL 200028
| 00131 | PCR machine | AppliedBiosystems | QuantStudio 1 |  | 2721224090023
| 64 | PCR machine | Roche | LightCycler 480 II, 96 |  | 32258
| 73 | Mini Vortexer | Heathrow Scientific | 120567 |  | HSA09547
| 74 | MiniSpin | Eppendorf | Various |  | 5452HR504062.
| 98 | Axygen Plate Centrifuge | Corning | Various |  | 22132009
| 89 | BioShaker iQ | Qinstrument | BioShaker iQ |  | 26656
Note: Throughout workbook, editable fields are denoted by yellow highlighting.

## Sample List
2025-11-05_095412_DEN078_RNaseP-16S.eds
| Sample Information |  | Workbook Assumed Sample Positions/Layout
| Sample No. | Sample Barcode/ID |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| 1 | PD005745RME2_D_Hu_01 | A | 1: Standard 1; 4: 1; 7: 9; 10: 17
| 2 | PD005599RME1_D_Hu_01 | B | 1: Standard 2; 4: 2; 7: 10; 10: 18
| 3 | PD005595RME1_D_Hu_01 | C | 1: Standard 3; 4: 3; 7: 11; 10: 19
| 4 | PD005590RME1_D_Hu_01 | D | 1: Standard 4; 4: 4; 7: 12; 10: 20
| 5 | PD005585RME1_D_Hu_01 | E | 1: Standard 5; 4: 5; 7: 13; 10: 21
| 6 | PD005594RME1_D_Hu_01 | F | 1: Standard 6; 4: 6; 7: 14; 10: 22
| 7 | PD005593RME1_D_Hu_01 | G | 1: POS Control; 4: 7; 7: 15; 10: 23
| 8 | PD005592RME1_D_Hu_01 | H | 1: NEG Control; 4: 8; 7: 16; 10: 24
| 9 | PD005584RME1_D_Hu_01
| 10 | PD005578RME1_D_Hu_01
| 11 | PD005575RME1_D_Hu_01
| 12 | PD005572RME1_D_Hu_01
| 13 | PD005571RME1_D_Hu_01
| 14 | PD005568RME1_D_Hu_01
| 15 | PD005565RME1_D_Hu_01
| 16 | PD005564RME4_D_Hu_01
| 17 | PD005563RME1_D_Hu_01
| 18 | PD005562RME1_D_Hu_01
| 19 | PD005561RME1_D_Hu_01
| 20 | PD005558RME1_D_Hu_01
| 21 | PD005557RME1_D_Hu_01
| 22 | PD005556RME1_D_Hu_01
| 23 | DEN078-POS_D_Hu_01
| 24 | DEN078-NEG_D_Hu_01
| POS Control | Human Male DNA 10ng/uL
| NEG Control | NF-Water

## Standards
2025-11-05_095412_DEN078_RNaseP-16S.eds
PCR Standard Curve Preparation
RNase P Standard Curve Preparation
| Human Genomic DNA Stock Concentration (µg/mL): |  |  | 209
| Calculations for Standard Curve Preparation: |  |  |  |  | Optimal (Theoretical) Standard Curve:
|  | Volumes Required (uL) |  |  | Final Concentration (pg/µl or ng/mL) |  | Final Concentration (pg/µL or ng/mL)
| Standard | Previous Dilution | Diluent (TE Buffer) | Total* |  | Standard
| S1 | 14.36 (Stock) | 5.64 | 20 | 150062 | S1 | 150000
| S2 | 4.50 (S1) | 15.5 | 20 | 33763.95 | S2 | 33750
| S3 | 4.50 (S2) | 15.5 | 20 | 7596.8888 | S3 | 7593.75
| S4 | 4.50 (S3) | 15.5 | 20 | 1709.3 | S4 | 1708.5938
| S5 | 4.50 (S4) | 15.5 | 20 | 384.5925 | S5 | 384.4336
| S6 | 4.50 (S5) | 15.5 | 20 | 86.5333 | S6 | 86.4976
*Prior to using 4.5uL to prepare the next sequential standard.
Note:
p20 pipette has mininum changeable unit of 0.02 leading to very small changes in concentration. This is factored in to produce above table.
16S Standard Curve Preparation
Dilute Microbial Community DNA Standard to a Concentration of 20000ng/mL Prior to Preparing Curve - This is added as the First Standard
Calculations for Standard Curve Preparation:
|  | Volumes Required (uL) |  |  | Theoretical Concentration (pg/µl)
| Standard | Previous Dilution | Diluent (TE Buffer) | Total*
| S1 | 16.00 (Stock) | 0 | 16 | 20000
| S2 | 2.00 (S1) | 14 | 16 | 2500
| S3 | 2.00 (S2) | 14 | 16 | 312.5
| S4 | 2.00 (S3) | 14 | 16 | 39.06
| S5 | 2.00 (S4) | 14 | 16 | 4.88
| S6 | 2.00 (S5) | 14 | 16 | 0.61
*Prior to using 2uL to prepare the next sequential standard.
Note:
For the NTC, add the diluent used for these standards

## qPCR Setup
2025-11-05_095412_DEN078_RNaseP-16S.eds
| Plate Layout for qPCR | [12 empty] | Master Mix Calculations
[14 empty] | Number of Samples: | 24
Standards Plated in Triplicate, Test Samples (1-24) Plated in Triplicate and Diluted 1 in 20:
[14 empty] | Component |  | Stock Conc. |  | Reaction Conc. |  | Per Well (uL) | Standards / Controls (uL) | Total Required* (uL)
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | [8 empty] | Required | (+10%)
| A | 1: S1; 4: 1 (1 in 20); 7: 9 (1 in 20); 10: 17 (1 in 20) | Fast Advanced Master Mix |  | 2x |  | 1x |  | 10 | 240 | 960 | 1056
| B | 1: S2; 4: 2 (1 in 20); 7: 10 (1 in 20); 10: 18 (1 in 20) | RNase P Primer-Probe (FAM) |  | 20x |  | 1x |  | 0.5 | 12 | 48 | 52.8
| C | 1: S3; 4: 3 (1 in 20); 7: 11 (1 in 20); 10: 19 (1 in 20) | 16S Primer-Probe (VIC) |  | 20x |  | 1x |  | 1 | 24 | 96 | 105.6
| D | 1: S4; 4: 4 (1 in 20); 7: 12 (1 in 20); 10: 20 (1 in 20) | TE Buffer/NF-Water |  | N/A |  | N/A |  | 4.5 | 108 | 432 | 475.2
| E | 1: S5; 4: 5 (1 in 20); 7: 13 (1 in 20); 10: 21 (1 in 20) | [5 empty] | Totals: | 16 | 384 | 1536 | 1689.6
| F | 1: S6; 4: 6 (1 in 20); 7: 14 (1 in 20); 10: 22 (1 in 20) | [6 empty] | *Total Required Including Standards
| G | 1: Human DNA Male - 10ng/µL; 4: 7 (1 in 20); 7: 15 (1 in 20); 10: 23 (1 in 20)
| H | 1: NTC; 4: 8 (1 in 20); 7: 16 (1 in 20); 10: 24 (1 in 20)
Cycling Conditions on PCR machine
[14 empty] | Step | Temperature |  | Time |  | # Cycles
[14 empty] | UNG Incubation | 50˚C |  | 2 Minutes |  | 1
[14 empty] | Enzyme Activation | 95˚C |  | 20 Seconds |  | 1
[14 empty] | Denaturation | 95˚C |  | 1 Second |  | 45
[14 empty] | Anneal/Extension | 60˚C |  | 20 Seconds

## LC Raw Results (hidden)
Select All in .txt Export File and Paste into Cell A2

## LC Calculations (hidden)
Plate Map | [12 empty] | Standards | [6 empty] | Options
PCR1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | Sample | Cp1 | Cp2 | Cp3 | Mean Cp | SEM | LC Conc. Mean | PCR2 | Calculated Mean Concs: | Conc. Used:
A | A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9 | A10 | A11 | A12 | Std. 1 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
B | B1 | B2 | B3 | B4 | B5 | B6 | B7 | B8 | B9 | B10 | B11 | B12 | Std. 2 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
C | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | C11 | C12 | Std. 3 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
D | D1 | D2 | D3 | D4 | D5 | D6 | D7 | D8 | D9 | D10 | D11 | D12 | Std. 4 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
E | E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9 | E10 | E11 | E12 | Std. 5 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
F | F1 | F2 | F3 | F4 | F5 | F6 | F7 | F8 | F9 | F10 | F11 | F12 | Std. 6 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
G | G1 | G2 | G3 | G4 | G5 | G6 | G7 | G8 | G9 | G10 | G11 | G12 | HuMale | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
H | H1 | H2 | H3 | H4 | H5 | H6 | H7 | H8 | H9 | H10 | H11 | H12 | NTC | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
Cps | [12 empty] | Samples
LC Concs?
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | Sample | Cp1 | Cp2 | Cp3 | Mean Cp | SEM | LC Conc. Mean | No | Calculated Mean Concs: | Conc. Used:
A | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 1 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
B | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 2 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
C | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 3 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
D | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 4 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
E | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 5 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
F | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 6 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
G | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 7 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
H | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 8 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
[13 empty] | 9 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
Concentrations | [12 empty] | 10 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
[13 empty] | 11 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 12 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
A | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 13 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
B | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 14 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
C | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 15 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
D | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 16 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
E | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 17 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
F | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 18 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
G | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 19 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
H | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | #VALUE! |  |  | 20 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
[13 empty] | 21 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
[13 empty] | 22 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
[13 empty] | 23 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0
[13 empty] | 24 | #VALUE! |  |  | 0 | 0 | #VALUE! |  | 0 | 0

## LC Analysis (hidden)
Paste Results File into Cell A2 of LC Raw Results
[13 empty] | Promega Conc. (ug/mL): | 244
|  |  | PCR1 or PCR2?: | PCR2 |  |  |  |  | Standard Curve Analysis |  |  |  | Promega Used (uL): | 12.3 |  |  |  | Pass/Fail?
Tabulated Averages
|  |  | PCR1 Means |  |  | PCR2 Means |  |  | Actual Results |  |  |  | Expected Results |  |  |  |  | #DIV/0!
| Sample No. | Sample ID | Cp | Conc. (ng/mL) | Cp SEM | 1/10 Cp | Conc. (ng/mL)* | Cp SEM | Standard | Mean Cp | Mean Conc. (ng/mL)* |  | Standard | Mean Cp* | Mean Conc. (ng/mL)
| 1 | PD005745RME2_D_Hu_01 |  |  |  | 0 | 0 | 0 | S1 | 0 | 0 |  | S1 | 22.8861 | 150060
| 2 | PD005599RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S2 | 0 | 0 |  | S2 | 25.0381 | 33763.5
| 3 | PD005595RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S3 | 0 | 0 |  | S3 | 27.1901 | 7596.7875
| 4 | PD005590RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S4 | 0 | 0 |  | S4 | 29.3421 | 1709.2772
| 5 | PD005585RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S5 | 0 | 0 |  | S5 | 31.4941 | 384.5874
| 6 | PD005594RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S6 | 0 | 0 |  | S6 | 33.6461 | 86.5322
| 7 | PD005593RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | POS | 0 | 0 |  | POS | 26.7936 | 10000
| 8 | PD005592RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | NTC | 0 | 0 |  | NTC | 0 | 0
| 9 | PD005584RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | *Based on the Following: |  |  |  | *Based on the Following:
| 10 | PD005578RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | Cp = ('Slope'*'LogConc')+'Y-Intercept' |  |  |  | -> Formula "Predicted Cp = LOG(((3.6*(10^11))/'InputGenomeCopies'),2)"
| 11 | PD005575RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | Conc (ng/mL) = 10^(('Mean Cp'-'Y-Intercept')/'Slope') |  |  |  | -> 'InputGenomeCopies' = [Sample in Reaction (uL)*Conc(ng/uL)]/[Genome Mass(ng)]
| 12 | PD005572RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 |  |  |  |  | -> 1 Genome = 0.00646246ng
| 13 | PD005571RME1_D_Hu_01 |  |  |  | 0 | 0 | 0
| 14 | PD005568RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | Standard Curve Breakdown
| 15 | PD005565RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | Actual Results | [5 empty] | Full Actual Results
| 16 | PD005564RME4_D_Hu_01 |  |  |  | 0 | 0 | 0 | Standard | Mean Cp | Mean Conc. (ng/mL) | Log Conc.* |  |  | Standard | Cp 1 | Cp 2 | Cp 3 | SEM | ΔCt
| 17 | PD005563RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S1 | 0 | 0 | 5.1763 |  |  | S1 | #VALUE! | 0 | 0 | 0 | -
| 18 | PD005562RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S2 | 0 | 0 | 4.5284 |  |  | S2 | #VALUE! | 0 | 0 | 0 | 0
| 19 | PD005561RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S3 | 0 | 0 | 3.8806 |  |  | S3 | #VALUE! | 0 | 0 | 0 | 0
| 20 | PD005558RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S4 | 0 | 0 | 3.2328 |  |  | S4 | #VALUE! | 0 | 0 | 0 | 0
| 21 | PD005557RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S5 | 0 | 0 | 2.585 |  |  | S5 | #VALUE! | 0 | 0 | 0 | 0
| 22 | PD005556RME1_D_Hu_01 |  |  |  | 0 | 0 | 0 | S6 | 0 | 0 | 1.9372 |  |  | S6 | #VALUE! | 0 | 0 | 0 | 0
| 23 | DEN078-POS_D_Hu_01 |  |  |  | 0 | 0 | 0 | *Based on Expected Concentration
| 24 | DEN078-NEG_D_Hu_01 |  |  |  | 0 | 0 | 0
| POS Control | Human Male DNA 10ng/uL |  |  |  | 0 | 0 | 0 | Graphical Representation
| NEG Control | NF-Water |  |  |  | 0 | 0 | 0
*Actual Sample Concentration, Calculated as 10x LightCycler Output
Conditional Formatting Key
| Text | Cp too Low/Conc too High, Requires Repeat with Dilution in PCR2 | [17 empty] | Plot Analysis:
| Text | Cp too Low/Conc too High, Result may be Unreliable | [17 empty] | 0 | = Y-intercept
| Text | Cp too High/Conc too Low to Accurately Quantify | [17 empty] | 0 | = Slope
| Text | SEM too High to be reliable, Investigate Individual Cps
[20 empty] | #DIV/0! | = PCR Efficiency
[20 empty] | #DIV/0! | = %PCR Efficiency
[20 empty] | #DIV/0! | = R Squared
Acceptance Criteria
[9 empty] | Criteria |  | Acceptance Range | Results | Pass/Fail?
[9 empty] | PCR Efficiency Result |  | 85%-115% | #DIV/0! | #DIV/0!
[9 empty] | R^2 |  | ≥0.98 | #DIV/0! | #DIV/0!
[9 empty] | NTC |  | Ct/Cp ≥ 38 | 0 | PASS
[9 empty] | Positive Control |  | 5ng/µL-15ng/µL | 0 | FAIL
[12 empty] | Passed Acceptance Criteria? = | #DIV/0!

## QS Raw Results
2025-11-05_095412_DEN078_RNaseP-16S.eds
* Block Type = 96-Well 0.2-mL Block
* Calibration Background is expired = No
* Calibration Background performed on = 11-18-2025
* Calibration Pure Dye ABY is expired = No
* Calibration Pure Dye ABY performed on = 11-18-2025
* Calibration Pure Dye FAM is expired = No
* Calibration Pure Dye FAM performed on = 11-18-2025
* Calibration Pure Dye JUN is expired = No
* Calibration Pure Dye JUN performed on = 11-18-2025
* Calibration Pure Dye ROX is expired = No
* Calibration Pure Dye ROX performed on = 11-18-2025
* Calibration Pure Dye SYBR is expired = No
* Calibration Pure Dye SYBR performed on = 11-18-2025
* Calibration Pure Dye VIC is expired = No
* Calibration Pure Dye VIC performed on = 11-18-2025
* Calibration ROI is expired = No
* Calibration ROI performed on = 11-18-2025
* Calibration Uniformity is expired = No
* Calibration Uniformity performed on = 11-18-2025
* Chemistry = TAQMAN
* Date Created = 2025-12-01 16:54:21 PM GMT
* Experiment Barcode = NA
* Experiment Comment = NA
* Experiment File Name = D:\Users\Administrator\experiments\TRIOMIC RNase P-16S duplex SOP-101 v1\DEN078\2025-11-05_095412_DEN078_RNaseP-16S.eds
* Experiment Name = 2025-11-05_095412_DEN078_RNaseP-16S
* Experiment Run End Time = 2025-12-01 17:27:18 PM GMT
* Experiment Type = Standard Curve
* Instrument Name = 2721224090023
* Instrument Serial Number = 2721224090023
* Instrument Type = QuantStudio™ 1 System
* Passive Reference = ROX
* Post-read Stage/Step =
* Pre-read Stage/Step =
* Quantification Cycle Method = Ct
* Signal Smoothing On = true
* Stage/ Cycle where Ct Analysis is performed = Stage2, Step2
* User Name = NA
[Results]
Well | Well Position | Omit | Sample Name | Target Name | Task | Reporter | Quencher | CT | Ct Mean | Ct SD | Quantity | Quantity Mean | Quantity SD | Y-Intercept | R(superscript 2) | Slope | Efficiency | Automatic Ct Threshold | Ct Threshold | Automatic Baseline | Baseline Start | Baseline End | Amp Status | Comments | Cq Conf | CQCONF | EXPFAIL | HIGHSD
1 | A1 | FALSE | Standard 1 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 21.844 | 21.936 | 0.134 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 17 | Amp |  | 0.972 | N | N | N
1 | A1 | FALSE | Standard 1 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 13.061 | 13.128 | 0.1 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 8 | Amp |  | 0.958 | N | N | N
2 | A2 | FALSE | Standard 1 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 21.874 | 21.936 | 0.134 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 17 | Amp |  | 0.958 | N | N | N
2 | A2 | FALSE | Standard 1 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 13.081 | 13.128 | 0.1 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 8 | Amp |  | 0.954 | N | N | N
3 | A3 | FALSE | Standard 1 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 22.089 | 21.936 | 0.134 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 17 | Amp |  | 0.954 | N | N | N
3 | A3 | FALSE | Standard 1 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 13.243 | 13.128 | 0.1 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 8 | Amp |  | 0.956 | N | N | N
4 | A4 | FALSE | Sample 1 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.4 | 32.741 | 0.302 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.95 | N | N | N
4 | A4 | FALSE | Sample 1 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.882 | 20.918 | 0.079 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.961 | N | N | N
5 | A5 | FALSE | Sample 1 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.847 | 32.741 | 0.302 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.941 | N | N | N
5 | A5 | FALSE | Sample 1 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 21.008 | 20.918 | 0.079 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.965 | N | N | N
6 | A6 | FALSE | Sample 1 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.975 | 32.741 | 0.302 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.95 | N | N | N
6 | A6 | FALSE | Sample 1 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.864 | 20.918 | 0.079 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.965 | N | N | N
7 | A7 | FALSE | Sample 9 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.863 | 26.892 | 0.04 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.942 | N | N | N
7 | A7 | FALSE | Sample 9 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 26.019 | 26.071 | 0.051 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.952 | N | N | N
8 | A8 | FALSE | Sample 9 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.876 | 26.892 | 0.04 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.945 | N | N | N
8 | A8 | FALSE | Sample 9 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 26.074 | 26.071 | 0.051 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.952 | N | N | N
9 | A9 | FALSE | Sample 9 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.938 | 26.892 | 0.04 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.941 | N | N | N
9 | A9 | FALSE | Sample 9 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 26.121 | 26.071 | 0.051 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.951 | N | N | N
10 | A10 | FALSE | Sample 17 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.993 | 32.193 | 0.177 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.945 | N | N | N
10 | A10 | FALSE | Sample 17 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.302 | 20.555 | 0.219 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.959 | N | N | N
11 | A11 | FALSE | Sample 17 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.257 | 32.193 | 0.177 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.953 | N | N | N
11 | A11 | FALSE | Sample 17 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.689 | 20.555 | 0.219 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.967 | N | N | N
12 | A12 | FALSE | Sample 17 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.329 | 32.193 | 0.177 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.947 | N | N | N
12 | A12 | FALSE | Sample 17 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.673 | 20.555 | 0.219 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.964 | N | N | N
13 | B1 | FALSE | Standard 2 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 24.036 | 24.007 | 0.031 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 19 | Amp |  | 0.959 | N | N | N
13 | B1 | FALSE | Standard 2 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.265 | 16.202 | 0.063 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.962 | N | N | N
14 | B2 | FALSE | Standard 2 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 24.011 | 24.007 | 0.031 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 19 | Amp |  | 0.949 | N | N | N
14 | B2 | FALSE | Standard 2 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.202 | 16.202 | 0.063 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.96 | N | N | N
15 | B3 | FALSE | Standard 2 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 23.974 | 24.007 | 0.031 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 19 | Amp |  | 0.944 | N | N | N
15 | B3 | FALSE | Standard 2 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.139 | 16.202 | 0.063 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.961 | N | N | N
16 | B4 | FALSE | Sample 2 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.687 | 31.814 | 0.112 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.957 | N | N | N
16 | B4 | FALSE | Sample 2 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.857 | 18.948 | 0.147 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.964 | N | N | N
17 | B5 | FALSE | Sample 2 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.851 | 31.814 | 0.112 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.951 | N | N | N
17 | B5 | FALSE | Sample 2 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.118 | 18.948 | 0.147 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.959 | N | N | N
18 | B6 | FALSE | Sample 2 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.903 | 31.814 | 0.112 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.956 | N | N | N
18 | B6 | FALSE | Sample 2 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.87 | 18.948 | 0.147 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 13 | Amp |  | 0.963 | N | N | N
19 | B7 | FALSE | Sample 10 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.749 | 29.66 | 0.079 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.951 | N | N | N
19 | B7 | FALSE | Sample 10 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 17.718 | 17.771 | 0.059 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 12 | Amp |  | 0.964 | N | N | N
20 | B8 | FALSE | Sample 10 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.599 | 29.66 | 0.079 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.963 | N | N | N
20 | B8 | FALSE | Sample 10 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 17.76 | 17.771 | 0.059 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 13 | Amp |  | 0.965 | N | N | N
21 | B9 | FALSE | Sample 10 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.633 | 29.66 | 0.079 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.953 | N | N | N
21 | B9 | FALSE | Sample 10 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 17.834 | 17.771 | 0.059 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 12 | Amp |  | 0.963 | N | N | N
22 | B10 | FALSE | Sample 18 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.398 | 32.458 | 0.122 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.95 | N | N | N
22 | B10 | FALSE | Sample 18 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.426 | 19.601 | 0.186 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.961 | N | N | N
23 | B11 | FALSE | Sample 18 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.599 | 32.458 | 0.122 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.955 | N | N | N
23 | B11 | FALSE | Sample 18 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.582 | 19.601 | 0.186 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.972 | N | N | N
24 | B12 | FALSE | Sample 18 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.378 | 32.458 | 0.122 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.951 | N | N | N
24 | B12 | FALSE | Sample 18 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.796 | 19.601 | 0.186 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.967 | N | N | N
25 | C1 | FALSE | Standard 3 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.28 | 26.229 | 0.069 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.961 | N | N | N
25 | C1 | FALSE | Standard 3 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.777 | 19.678 | 0.091 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.964 | N | N | N
26 | C2 | FALSE | Standard 3 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.151 | 26.229 | 0.069 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 21 | Amp |  | 0.951 | N | N | N
26 | C2 | FALSE | Standard 3 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.597 | 19.678 | 0.091 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.963 | N | N | N
27 | C3 | FALSE | Standard 3 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.256 | 26.229 | 0.069 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.953 | N | N | N
27 | C3 | FALSE | Standard 3 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.662 | 19.678 | 0.091 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.962 | N | N | N
28 | C4 | FALSE | Sample 3 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.847 | 30.164 | 0.304 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.944 | N | N | N
28 | C4 | FALSE | Sample 3 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 25.404 | 25.872 | 0.423 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 20 | Amp |  | 0.948 | N | N | N
29 | C5 | FALSE | Sample 3 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.192 | 30.164 | 0.304 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.946 | N | N | N
29 | C5 | FALSE | Sample 3 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 25.987 | 25.872 | 0.423 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.957 | N | N | N
30 | C6 | FALSE | Sample 3 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.453 | 30.164 | 0.304 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.947 | N | N | N
30 | C6 | FALSE | Sample 3 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 26.226 | 25.872 | 0.423 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.955 | N | N | N
31 | C7 | FALSE | Sample 11 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.728 | 29.691 | 0.048 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.947 | N | N | N
31 | C7 | FALSE | Sample 11 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.974 | 19.038 | 0.068 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.962 | N | N | N
32 | C8 | FALSE | Sample 11 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.636 | 29.691 | 0.048 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.953 | N | N | N
32 | C8 | FALSE | Sample 11 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.029 | 19.038 | 0.068 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.96 | N | N | N
33 | C9 | FALSE | Sample 11 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.709 | 29.691 | 0.048 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.951 | N | N | N
33 | C9 | FALSE | Sample 11 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.11 | 19.038 | 0.068 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.96 | N | N | N
34 | C10 | FALSE | Sample 19 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.911 | 29.765 | 0.126 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.953 | N | N | N
34 | C10 | FALSE | Sample 19 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.99 | 18.982 | 0.011 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.965 | N | N | N
35 | C11 | FALSE | Sample 19 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.697 | 29.765 | 0.126 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.954 | N | N | N
35 | C11 | FALSE | Sample 19 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.988 | 18.982 | 0.011 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.964 | N | N | N
36 | C12 | FALSE | Sample 19 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.688 | 29.765 | 0.126 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.955 | N | N | N
36 | C12 | FALSE | Sample 19 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.97 | 18.982 | 0.011 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.968 | N | N | N
37 | D1 | FALSE | Standard 4 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.527 | 28.544 | 0.114 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.953 | N | N | N
37 | D1 | FALSE | Standard 4 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 22.983 | 22.935 | 0.08 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 18 | Amp |  | 0.965 | N | N | N
38 | D2 | FALSE | Standard 4 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.665 | 28.544 | 0.114 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.944 | N | N | N
38 | D2 | FALSE | Standard 4 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 22.979 | 22.935 | 0.08 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 18 | Amp |  | 0.963 | N | N | N
39 | D3 | FALSE | Standard 4 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.44 | 28.544 | 0.114 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.947 | N | N | N
39 | D3 | FALSE | Standard 4 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 22.843 | 22.935 | 0.08 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 17 | Amp |  | 0.962 | N | N | N
40 | D4 | FALSE | Sample 4 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.895 | 29.844 | 0.132 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.951 | N | N | N
40 | D4 | FALSE | Sample 4 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.101 | 19.065 | 0.085 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.959 | N | N | N
41 | D5 | FALSE | Sample 4 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.943 | 29.844 | 0.132 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.953 | N | N | N
41 | D5 | FALSE | Sample 4 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.127 | 19.065 | 0.085 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.964 | N | N | N
42 | D6 | FALSE | Sample 4 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.695 | 29.844 | 0.132 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.951 | N | N | N
42 | D6 | FALSE | Sample 4 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.968 | 19.065 | 0.085 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.96 | N | N | N
43 | D7 | FALSE | Sample 12 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.913 | 28.951 | 0.066 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.96 | N | N | N
43 | D7 | FALSE | Sample 12 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 15.842 | 15.949 | 0.098 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 10 | Amp |  | 0.965 | N | N | N
44 | D8 | FALSE | Sample 12 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.912 | 28.951 | 0.066 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.961 | N | N | N
44 | D8 | FALSE | Sample 12 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 15.973 | 15.949 | 0.098 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.963 | N | N | N
45 | D9 | FALSE | Sample 12 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 29.027 | 28.951 | 0.066 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.957 | N | N | N
45 | D9 | FALSE | Sample 12 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.033 | 15.949 | 0.098 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.958 | N | N | N
46 | D10 | FALSE | Sample 20 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.241 | 30.533 | 0.269 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.952 | N | N | N
46 | D10 | FALSE | Sample 20 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.685 | 21.274 | 0.565 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.963 | N | N | Y
47 | D11 | FALSE | Sample 20 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.588 | 30.533 | 0.269 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.953 | N | N | N
47 | D11 | FALSE | Sample 20 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 21.326 | 21.274 | 0.565 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.965 | N | N | Y
48 | D12 | FALSE | Sample 20 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.771 | 30.533 | 0.269 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.953 | N | N | N
48 | D12 | FALSE | Sample 20 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 21.812 | 21.274 | 0.565 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 17 | Amp |  | 0.969 | N | N | Y
49 | E1 | FALSE | Standard 5 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.902 | 30.929 | 0.026 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.957 | N | N | N
49 | E1 | FALSE | Standard 5 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 26.128 | 26.087 | 0.077 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.962 | N | N | N
50 | E2 | FALSE | Standard 5 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.93 | 30.929 | 0.026 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.946 | N | N | N
50 | E2 | FALSE | Standard 5 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 26.134 | 26.087 | 0.077 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 21 | Amp |  | 0.958 | N | N | N
51 | E3 | FALSE | Standard 5 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.954 | 30.929 | 0.026 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.944 | N | N | N
51 | E3 | FALSE | Standard 5 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 25.998 | 26.087 | 0.077 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 20 | Amp |  | 0.959 | N | N | N
52 | E4 | FALSE | Sample 5 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.807 | 28.475 | 0.29 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.95 | N | N | N
52 | E4 | FALSE | Sample 5 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 21.366 | 21.037 | 0.285 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.963 | N | N | N
53 | E5 | FALSE | Sample 5 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.35 | 28.475 | 0.29 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.952 | N | N | N
53 | E5 | FALSE | Sample 5 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.878 | 21.037 | 0.285 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.965 | N | N | N
54 | E6 | FALSE | Sample 5 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 28.268 | 28.475 | 0.29 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 24 | Amp |  | 0.944 | N | N | N
54 | E6 | FALSE | Sample 5 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.867 | 21.037 | 0.285 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.961 | N | N | N
55 | E7 | FALSE | Sample 13 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.739 | 31.818 | 0.162 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.954 | N | N | N
55 | E7 | FALSE | Sample 13 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.406 | 19.516 | 0.096 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.965 | N | N | N
56 | E8 | FALSE | Sample 13 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.712 | 31.818 | 0.162 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.954 | N | N | N
56 | E8 | FALSE | Sample 13 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.577 | 19.516 | 0.096 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.967 | N | N | N
57 | E9 | FALSE | Sample 13 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.004 | 31.818 | 0.162 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.949 | N | N | N
57 | E9 | FALSE | Sample 13 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.564 | 19.516 | 0.096 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.96 | N | N | N
58 | E10 | FALSE | Sample 21 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.365 | 30.654 | 0.296 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.954 | N | N | N
58 | E10 | FALSE | Sample 21 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.058 | 19.471 | 0.424 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.963 | N | N | N
59 | E11 | FALSE | Sample 21 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.64 | 30.654 | 0.296 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.948 | N | N | N
59 | E11 | FALSE | Sample 21 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.449 | 19.471 | 0.424 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 14 | Amp |  | 0.963 | N | N | N
60 | E12 | FALSE | Sample 21 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.956 | 30.654 | 0.296 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.954 | N | N | N
60 | E12 | FALSE | Sample 21 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 19.905 | 19.471 | 0.424 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.967 | N | N | N
61 | F1 | FALSE | Standard 6 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 33.328 | 33.204 | 0.174 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.909 | N | N | N
61 | F1 | FALSE | Standard 6 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.436 | 29.38 | 0.142 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.951 | N | N | N
62 | F2 | FALSE | Standard 6 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 33.005 | 33.204 | 0.174 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.946 | N | N | N
62 | F2 | FALSE | Standard 6 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.219 | 29.38 | 0.142 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.957 | N | N | N
63 | F3 | FALSE | Standard 6 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 33.279 | 33.204 | 0.174 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.944 | N | N | N
63 | F3 | FALSE | Standard 6 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.486 | 29.38 | 0.142 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.956 | N | N | N
64 | F4 | FALSE | Sample 6 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.825 | 30.69 | 0.12 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.949 | N | N | N
64 | F4 | FALSE | Sample 6 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 25.085 | 25.004 | 0.071 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 20 | Amp |  | 0.965 | N | N | N
65 | F5 | FALSE | Sample 6 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.595 | 30.69 | 0.12 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.951 | N | N | N
65 | F5 | FALSE | Sample 6 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 24.973 | 25.004 | 0.071 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 20 | Amp |  | 0.964 | N | N | N
66 | F6 | FALSE | Sample 6 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.649 | 30.69 | 0.12 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.947 | N | N | N
66 | F6 | FALSE | Sample 6 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 24.954 | 25.004 | 0.071 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 19 | Amp |  | 0.965 | N | N | N
67 | F7 | FALSE | Sample 14 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 33.041 | 33.237 | 0.332 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.951 | N | N | N
67 | F7 | FALSE | Sample 14 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.333 | 20.345 | 0.01 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.966 | N | N | N
68 | F8 | FALSE | Sample 14 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 33.619 | 33.237 | 0.332 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.952 | N | N | N
68 | F8 | FALSE | Sample 14 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.35 | 20.345 | 0.01 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.968 | N | N | N
69 | F9 | FALSE | Sample 14 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 33.05 | 33.237 | 0.332 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.978 | N | N | N
69 | F9 | FALSE | Sample 14 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.351 | 20.345 | 0.01 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.966 | N | N | N
70 | F10 | FALSE | Sample 22 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.266 | 32.178 | 0.149 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.953 | N | N | N
70 | F10 | FALSE | Sample 22 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.561 | 20.676 | 0.217 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.968 | N | N | N
71 | F11 | FALSE | Sample 22 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.006 | 32.178 | 0.149 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.953 | N | N | N
71 | F11 | FALSE | Sample 22 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.54 | 20.676 | 0.217 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 15 | Amp |  | 0.964 | N | N | N
72 | F12 | FALSE | Sample 22 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.262 | 32.178 | 0.149 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.954 | N | N | N
72 | F12 | FALSE | Sample 22 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 20.926 | 20.676 | 0.217 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 16 | Amp |  | 0.965 | N | N | N
73 | G1 | FALSE | Human DNA Male | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.927 | 27.154 | 0.596 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.945 | N | N | Y
74 | G2 | FALSE | Human DNA Male | Target 1 | UNKNOWN | FAM | NFQ-MGB | 26.705 | 27.154 | 0.596 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 21 | Amp |  | 0.933 | N | N | Y
75 | G3 | FALSE | Human DNA Male | Target 1 | UNKNOWN | FAM | NFQ-MGB | [10 empty] | TRUE | 0.207 | TRUE | 3 | 22 | Amp |  | 0.942 | N | N | Y
76 | G4 | FALSE | Sample 7 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 34.516 | 33.127 | 1.213 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 29 | Amp |  | 0.938 | N | N | Y
76 | G4 | FALSE | Sample 7 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.567 | 18.201 | 0.317 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 13 | Amp |  | 0.963 | N | N | N
77 | G5 | FALSE | Sample 7 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.584 | 33.127 | 1.213 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 28 | Amp |  | 0.957 | N | N | Y
77 | G5 | FALSE | Sample 7 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.014 | 18.201 | 0.317 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 13 | Amp |  | 0.964 | N | N | N
78 | G6 | FALSE | Sample 7 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 32.28 | 33.127 | 1.213 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 27 | Amp |  | 0.953 | N | N | Y
78 | G6 | FALSE | Sample 7 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 18.022 | 18.201 | 0.317 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 13 | Amp |  | 0.963 | N | N | N
79 | G7 | FALSE | Sample 15 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 34.283 | 34.29 | 0.131 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 29 | Amp |  | 0.945 | N | N | N
79 | G7 | FALSE | Sample 15 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 22.09 | 22.122 | 0.028 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 17 | Amp |  | 0.963 | N | N | N
80 | G8 | FALSE | Sample 15 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 34.423 | 34.29 | 0.131 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 29 | Amp |  | 0.941 | N | N | N
80 | G8 | FALSE | Sample 15 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 22.135 | 22.122 | 0.028 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 17 | Amp |  | 0.964 | N | N | N
81 | G9 | FALSE | Sample 15 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 34.162 | 34.29 | 0.131 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 29 | Amp |  | 0.943 | N | N | N
81 | G9 | FALSE | Sample 15 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 22.142 | 22.122 | 0.028 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 17 | Amp |  | 0.962 | N | N | N
82 | G10 | FALSE | Sample 23 | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
82 | G10 | FALSE | Sample 23 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.566 | 29.804 | 0.3 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.951 | N | N | N
83 | G11 | FALSE | Sample 23 | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
83 | G11 | FALSE | Sample 23 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.703 | 29.804 | 0.3 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.954 | N | N | N
84 | G12 | FALSE | Sample 23 | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
84 | G12 | FALSE | Sample 23 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 30.141 | 29.804 | 0.3 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 25 | Amp |  | 0.954 | N | N | N
85 | H1 | FALSE | NTC | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
85 | H1 | FALSE | NTC | Target 2 | UNKNOWN | VIC | NFQ-MGB | 35.605 | 35.291 | 0.297 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 30 | Amp |  | 0.915 | N | N | N
86 | H2 | FALSE | NTC | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
86 | H2 | FALSE | NTC | Target 2 | UNKNOWN | VIC | NFQ-MGB | 35.014 | 35.291 | 0.297 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 30 | Amp |  | 0.927 | N | N | N
87 | H3 | FALSE | NTC | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
87 | H3 | FALSE | NTC | Target 2 | UNKNOWN | VIC | NFQ-MGB | 35.254 | 35.291 | 0.297 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 30 | Amp |  | 0.94 | N | N | N
88 | H4 | FALSE | Sample 8 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.172 | 30.155 | 0.085 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.942 | N | N | N
88 | H4 | FALSE | Sample 8 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 17.816 | 17.745 | 0.073 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 13 | Amp |  | 0.962 | N | N | N
89 | H5 | FALSE | Sample 8 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.23 | 30.155 | 0.085 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.935 | N | N | N
89 | H5 | FALSE | Sample 8 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 17.748 | 17.745 | 0.073 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 12 | Amp |  | 0.955 | N | N | N
90 | H6 | FALSE | Sample 8 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 30.062 | 30.155 | 0.085 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 25 | Amp |  | 0.956 | N | N | N
90 | H6 | FALSE | Sample 8 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 17.67 | 17.745 | 0.073 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 12 | Amp |  | 0.965 | N | N | N
91 | H7 | FALSE | Sample 16 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.149 | 31.246 | 0.085 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.938 | N | N | N
91 | H7 | FALSE | Sample 16 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.093 | 16.218 | 0.124 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.955 | N | N | N
92 | H8 | FALSE | Sample 16 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.308 | 31.246 | 0.085 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.934 | N | N | N
92 | H8 | FALSE | Sample 16 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.22 | 16.218 | 0.124 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.971 | N | N | N
93 | H9 | FALSE | Sample 16 | Target 1 | UNKNOWN | FAM | NFQ-MGB | 31.281 | 31.246 | 0.085 | [7 empty] | TRUE | 0.207 | TRUE | 3 | 26 | Amp |  | 0.957 | N | N | N
93 | H9 | FALSE | Sample 16 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 16.34 | 16.218 | 0.124 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 11 | Amp |  | 0.961 | N | N | N
94 | H10 | FALSE | Sample 24 | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
94 | H10 | FALSE | Sample 24 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.429 | 29.738 | 0.481 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.951 | N | N | N
95 | H11 | FALSE | Sample 24 | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
95 | H11 | FALSE | Sample 24 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 30.292 | 29.738 | 0.481 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 25 | Amp |  | 0.942 | N | N | N
96 | H12 | FALSE | Sample 24 | Target 1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.207 | TRUE | 3 | 44 | No Amp |  | 0 | Y | Y | N
96 | H12 | FALSE | Sample 24 | Target 2 | UNKNOWN | VIC | NFQ-MGB | 29.492 | 29.738 | 0.481 | [7 empty] | TRUE | 0.164 | TRUE | 3 | 24 | Amp |  | 0.957 | N | N | N

## QS Calculations (hidden)
Cps FAM | [12 empty] | Standards FAM | [6 empty] | Standards VIC
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | Sample | Cp1 | Cp2 | Cp3 | Mean Cp | SEM | Calculated Mean Conc. (ng/mL) | Sample | Cp1 | Cp2 | Cp3 | Mean Cp | SEM | Calculated Mean Conc. (ng/mL)
A | 21.844 | 21.874 | 22.089 | 32.4 | 32.847 | 32.975 | 26.863 | 26.876 | 26.938 | 31.993 | 32.257 | 32.329 | Std. 1 | 21.844 | 21.874 | 22.089 | 21.9357 | 0.062996 | 137416.3906 | Std. 1 | 13.061 | 13.081 | 13.243 | 13.1283 | 0.047049 | 19397.465
B | 24.036 | 24.011 | 23.974 | 31.687 | 31.851 | 31.903 | 29.749 | 29.599 | 29.633 | 32.398 | 32.599 | 32.378 | Std. 2 | 24.036 | 24.011 | 23.974 | 24.007 | 0.014704 | 35213.2489 | Std. 2 | 16.265 | 16.202 | 16.139 | 16.202 | 0.029698 | 2733.6508
C | 26.28 | 26.151 | 26.256 | 29.847 | 30.192 | 30.453 | 29.728 | 29.636 | 29.709 | 29.911 | 29.697 | 29.688 | Std. 3 | 26.28 | 26.151 | 26.256 | 26.229 | 0.032342 | 8172.6093 | Std. 3 | 19.777 | 19.597 | 19.662 | 19.6787 | 0.042969 | 297.9635
D | 28.527 | 28.665 | 28.44 | 29.895 | 29.943 | 29.695 | 28.913 | 28.912 | 29.027 | 30.241 | 30.588 | 30.771 | Std. 4 | 28.527 | 28.665 | 28.44 | 28.544 | 0.053485 | 1784.2897 | Std. 4 | 22.983 | 22.979 | 22.843 | 22.935 | 0.037571 | 37.3754
E | 30.902 | 30.93 | 30.954 | 28.807 | 28.35 | 28.268 | 31.739 | 31.712 | 32.004 | 30.365 | 30.64 | 30.956 | Std. 5 | 30.902 | 30.93 | 30.954 | 30.9287 | 0.012269 | 372.1185 | Std. 5 | 26.128 | 26.134 | 25.998 | 26.0867 | 0.036226 | 5.0117
F | 33.328 | 33.005 | 33.279 | 30.825 | 30.595 | 30.649 | 33.041 | 33.619 | 33.05 | 32.266 | 32.006 | 32.262 | Std. 6 | 33.328 | 33.005 | 33.279 | 33.204 | 0.082058 | 83.3893 | Std. 6 | 29.436 | 29.219 | 29.486 | 29.3803 | 0.06691 | 0.61387
G | 26.927 | 26.705 | 0 | 34.516 | 32.584 | 32.28 | 34.283 | 34.423 | 34.162 | 0 | 0 | 0 | HuMale | 26.927 | 26.705 | 0 | 26.816 | 7.2986 | 5556.2484 | N/A | #N/A | #N/A | #N/A | 0 | 0 | 0
H | 0 | 0 | 0 | 30.172 | 30.23 | 30.062 | 31.149 | 31.308 | 31.281 | 0 | 0 | 0 | NTC | 0 | 0 | 0 | 0 | 0 | 0 | NTC | 35.605 | 35.014 | 35.254 | 35.291 | 0.14012 | 0.014177
[13 empty] | Samples FAM | [6 empty] | Samples VIC
Cps VIC
[13 empty] | Sample | Cp1 | Cp2 | Cp3 | Mean Cp | SEM | Calculated Mean Conc. (ng/mL) | Sample | Cp1 | Cp2 | Cp3 | Mean Cp | SEM | Calculated Mean Conc. (ng/mL)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 32.4 | 32.847 | 32.975 | 32.7407 | 0.14231 | 113.0796 | 1 | 20.882 | 21.008 | 20.864 | 20.918 | 0.036986 | 135.2168
A | 13.061 | 13.081 | 13.243 | 20.882 | 21.008 | 20.864 | 26.019 | 26.074 | 26.121 | 20.302 | 20.689 | 20.673 | 2 | 31.687 | 31.851 | 31.903 | 31.8137 | 0.053144 | 207.9831 | 2 | 18.857 | 19.118 | 18.87 | 18.9483 | 0.069334 | 474.6456
B | 16.265 | 16.202 | 16.139 | 18.857 | 19.118 | 18.87 | 17.718 | 17.76 | 17.834 | 19.426 | 19.582 | 19.796 | 3 | 29.847 | 30.192 | 30.453 | 30.164 | 0.14329 | 615.1505 | 3 | 25.404 | 25.987 | 26.226 | 25.8723 | 0.19932 | 5.7455
C | 19.777 | 19.597 | 19.662 | 25.404 | 25.987 | 26.226 | 18.974 | 19.029 | 19.11 | 18.99 | 18.988 | 18.97 | 4 | 29.895 | 29.943 | 29.695 | 29.8443 | 0.062006 | 758.9988 | 4 | 19.101 | 19.127 | 18.968 | 19.0653 | 0.040206 | 440.5303
D | 22.983 | 22.979 | 22.843 | 19.101 | 19.127 | 18.968 | 15.842 | 15.973 | 16.033 | 20.685 | 21.326 | 21.812 | 5 | 28.807 | 28.35 | 28.268 | 28.475 | 0.13691 | 1867.0836 | 5 | 21.366 | 20.878 | 20.867 | 21.037 | 0.13434 | 125.3381
E | 26.128 | 26.134 | 25.998 | 21.366 | 20.878 | 20.867 | 19.406 | 19.577 | 19.564 | 19.058 | 19.449 | 19.905 | 6 | 30.825 | 30.595 | 30.649 | 30.6897 | 0.056697 | 435.4236 | 6 | 25.085 | 24.973 | 24.954 | 25.004 | 0.03337 | 9.9941
F | 29.436 | 29.219 | 29.486 | 25.085 | 24.973 | 24.954 | 20.333 | 20.35 | 20.351 | 20.561 | 20.54 | 20.926 | 7 | 34.516 | 32.584 | 32.28 | 33.1267 | 0.5717 | 87.738 | 7 | 18.567 | 18.014 | 18.022 | 18.201 | 0.14943 | 764.333
G | #N/A | #N/A | #N/A | 18.567 | 18.014 | 18.022 | 22.09 | 22.135 | 22.142 | 29.566 | 29.703 | 30.141 | 8 | 30.172 | 30.23 | 30.062 | 30.1547 | 0.040225 | 618.9362 | 8 | 17.816 | 17.748 | 17.67 | 17.7447 | 0.034439 | 1022.4145
H | 35.605 | 35.014 | 35.254 | 17.816 | 17.748 | 17.67 | 16.093 | 16.22 | 16.34 | 29.429 | 30.292 | 29.492 | 9 | 26.863 | 26.876 | 26.938 | 26.8923 | 0.018893 | 5284.3273 | 9 | 26.019 | 26.074 | 26.121 | 26.0713 | 0.024066 | 5.061
[13 empty] | 10 | 29.749 | 29.599 | 29.633 | 29.6603 | 0.037074 | 856.5843 | 10 | 17.718 | 17.76 | 17.834 | 17.7707 | 0.027686 | 1005.6073
[13 empty] | 11 | 29.728 | 29.636 | 29.709 | 29.691 | 0.022896 | 839.4895 | 11 | 18.974 | 19.029 | 19.11 | 19.0377 | 0.03225 | 448.3693
[13 empty] | 12 | 28.913 | 28.912 | 29.027 | 28.9507 | 0.031164 | 1365.7417 | 12 | 15.842 | 15.973 | 16.033 | 15.9493 | 0.046044 | 3211.4304
[13 empty] | 13 | 31.739 | 31.712 | 32.004 | 31.8183 | 0.076065 | 207.346 | 13 | 19.406 | 19.577 | 19.564 | 19.5157 | 0.044876 | 330.5922
[13 empty] | 14 | 33.041 | 33.619 | 33.05 | 33.2367 | 0.1561 | 81.6177 | 14 | 20.333 | 20.35 | 20.351 | 20.3447 | 0.0047687 | 194.8806
[13 empty] | 15 | 34.283 | 34.423 | 34.162 | 34.2893 | 0.061573 | 40.8568 | 15 | 22.09 | 22.135 | 22.142 | 22.1223 | 0.013303 | 62.7463
[13 empty] | 16 | 31.149 | 31.308 | 31.281 | 31.246 | 0.040108 | 302.0562 | 16 | 16.093 | 16.22 | 16.34 | 16.2177 | 0.058226 | 2706.4838
[13 empty] | 17 | 31.993 | 32.257 | 32.329 | 32.193 | 0.083395 | 162.0819 | 17 | 20.302 | 20.689 | 20.673 | 20.5547 | 0.10322 | 170.4614
[13 empty] | 18 | 32.398 | 32.599 | 32.378 | 32.4583 | 0.05762 | 136.1401 | 18 | 19.426 | 19.582 | 19.796 | 19.6013 | 0.087566 | 313.0215
[13 empty] | 19 | 29.911 | 29.697 | 29.688 | 29.7653 | 0.059506 | 799.4555 | 19 | 18.99 | 18.988 | 18.97 | 18.9827 | 0.0051926 | 464.3694
[13 empty] | 20 | 30.241 | 30.588 | 30.771 | 30.5333 | 0.1269 | 482.5503 | 20 | 20.685 | 21.326 | 21.812 | 21.2743 | 0.26647 | 107.739
[13 empty] | 21 | 30.365 | 30.64 | 30.956 | 30.6537 | 0.13941 | 445.8506 | 21 | 19.058 | 19.449 | 19.905 | 19.4707 | 0.19984 | 340.2136
[13 empty] | 22 | 32.266 | 32.006 | 32.262 | 32.178 | 0.070225 | 163.6879 | 22 | 20.561 | 20.54 | 20.926 | 20.6757 | 0.10232 | 157.8065
[13 empty] | 23 | 0 | 0 | 0 | 0 | 0 | 0 | 23 | 29.566 | 29.703 | 30.141 | 29.8033 | 0.14158 | 0.46877
[13 empty] | 24 | 0 | 0 | 0 | 0 | 0 | 0 | 24 | 29.429 | 30.292 | 29.492 | 29.7377 | 0.22679 | 0.48881

## QS RNase P Analysis
2025-11-05_095412_DEN078_RNaseP-16S.eds
[10 empty] | Promega Conc. (ug/mL): | 209
[6 empty] | Standard Curve Analysis |  |  |  | Promega Used (uL): | 14.36 |  |  |  | Pass/Fail?
Tabulated Averages
[6 empty] | Actual Results |  |  |  | Expected Results |  |  |  |  | PASSED
| Sample No. | Sample ID | 1/10 Cp | Conc. (ng/μL)* | Cp SEM | Standard | Mean Cp | Mean Conc. (ng/mL)* |  | Standard | Mean Cp* | Mean Conc. (ng/mL)
| 1 | PD005745RME2_D_Hu_01 | 32.7407 | 1.1308 | 0.14231 | S1 | 21.9357 | 137416.3906 |  | S1 | 22.8861 | 150062
| 2 | PD005599RME1_D_Hu_01 | 31.8137 | 2.0798 | 0.053144 | S2 | 24.007 | 35213.2489 |  | S2 | 25.0381 | 33763.95
| 3 | PD005595RME1_D_Hu_01 | 30.164 | 6.1515 | 0.14329 | S3 | 26.229 | 8172.6093 |  | S3 | 27.1901 | 7596.8888
| 4 | PD005590RME1_D_Hu_01 | 29.8443 | 7.59 | 0.062006 | S4 | 28.544 | 1784.2897 |  | S4 | 29.3421 | 1709.3
| 5 | PD005585RME1_D_Hu_01 | 28.475 | 18.6708 | 0.13691 | S5 | 30.9287 | 372.1185 |  | S5 | 31.4941 | 384.5925
| 6 | PD005594RME1_D_Hu_01 | 30.6897 | 4.3542 | 0.056697 | S6 | 33.204 | 83.3893 |  | S6 | 33.6461 | 86.5333
| 7 | PD005593RME1_D_Hu_01 | 33.1267 | 0.87738 | 0.5717 | POS | 26.816 | 5556.2484 |  | POS | 26.7936 | 10000
| 8 | PD005592RME1_D_Hu_01 | 30.1547 | 6.1894 | 0.040225 | NTC | 0 | 0 |  | NTC | 0 | 0
| 9 | PD005584RME1_D_Hu_01 | 26.8923 | 52.8433 | 0.018893 | *Based on the Following: |  |  |  | *Based on the Following:
| 10 | PD005578RME1_D_Hu_01 | 29.6603 | 8.5658 | 0.037074 | Cp = ('Slope'*'LogConc')+'Y-Intercept' |  |  |  | -> Formula "Predicted Cp = LOG(((3.6*(10^11))/'InputGenomeCopies'),2)"
| 11 | PD005575RME1_D_Hu_01 | 29.691 | 8.3949 | 0.022896 | Conc (ng/mL) = 10^(('Mean Cp'-'Y-Intercept')/'Slope') |  |  |  | -> 'InputGenomeCopies' = [Sample in Reaction (uL)*Conc(ng/uL)]/[Genome Mass(ng)]
| 12 | PD005572RME1_D_Hu_01 | 28.9507 | 13.6574 | 0.031164 |  |  |  |  | -> 1 Genome = 0.00646246ng
| 13 | PD005571RME1_D_Hu_01 | 31.8183 | 2.0735 | 0.076065
| 14 | PD005568RME1_D_Hu_01 | 33.2367 | 0.81618 | 0.1561 | Standard Curve Breakdown
| 15 | PD005565RME1_D_Hu_01 | 34.2893 | 0.40857 | 0.061573 | Actual Results | [5 empty] | Full Actual Results
| 16 | PD005564RME4_D_Hu_01 | 31.246 | 3.0206 | 0.040108 | Standard | Mean Cp | Mean Conc. (ng/mL) | Log Conc.* |  |  | Standard | Cp 1 | Cp 2 | Cp 3 | SEM | ΔCt
| 17 | PD005563RME1_D_Hu_01 | 32.193 | 1.6208 | 0.083395 | S1 | 21.9357 | 137416.3906 | 5.1763 |  |  | S1 | 21.844 | 21.874 | 22.089 | 0.062996 | -
| 18 | PD005562RME1_D_Hu_01 | 32.4583 | 1.3614 | 0.05762 | S2 | 24.007 | 35213.2489 | 4.5285 |  |  | S2 | 24.036 | 24.011 | 23.974 | 0.014704 | 2.0713
| 19 | PD005561RME1_D_Hu_01 | 29.7653 | 7.9946 | 0.059506 | S3 | 26.229 | 8172.6093 | 3.8806 |  |  | S3 | 26.28 | 26.151 | 26.256 | 0.032342 | 2.222
| 20 | PD005558RME1_D_Hu_01 | 30.5333 | 4.8255 | 0.1269 | S4 | 28.544 | 1784.2897 | 3.2328 |  |  | S4 | 28.527 | 28.665 | 28.44 | 0.053485 | 2.315
| 21 | PD005557RME1_D_Hu_01 | 30.6537 | 4.4585 | 0.13941 | S5 | 30.9287 | 372.1185 | 2.585 |  |  | S5 | 30.902 | 30.93 | 30.954 | 0.012269 | 2.3847
| 22 | PD005556RME1_D_Hu_01 | 32.178 | 1.6369 | 0.070225 | S6 | 33.204 | 83.3893 | 1.9372 |  |  | S6 | 33.328 | 33.005 | 33.279 | 0.082058 | 2.2753
| 23 | DEN078-POS_D_Hu_01 | 0 | 0 | 0 | *Based on Expected Concentration
| 24 | DEN078-NEG_D_Hu_01 | 0 | 0 | 0
| POS Control | Human Male DNA 10ng/uL | 26.816 | 5.5562 | 7.2986 | Graphical Representation
| NEG Control | NF-Water | 0 | 0 | 0
| *Actual Sample Concentration, Calculated as 10x LightCycler Output | [14 empty] | Plot Analysis
[16 empty] | 39.9333 | = Y-intercept
| Conditional Formatting Key | [14 empty] | -3.5028 | = Slope
| Text | Cp too Low/Conc too High, Result may be Unreliable
| Text | Cp too High/Conc too Low to Accurately Quantify | [13 empty] | 1.9297 | = PCR Efficiency
| Text | SEM too High to be reliable, Investigate Individual Cps | [13 empty] | 92.9674 | = %PCR Efficiency
[16 empty] | 0.99951 | = R Squared
Acceptance Criteria
[6 empty] | Criteria | Acceptance Range | Results | Pass/Fail?
[6 empty] | PCR Efficiency Result | 85%-115% | 92.9674 | PASS
[6 empty] | R^2 | ≥0.98 | 0.99951 | PASS
[6 empty] | NTC | Ct/Cp ≥ 38 | 0 | PASS
[6 empty] | Positive Control | 5ng/µL-15ng/µL | 5.5562 | PASS
[6 empty] | Y-intercept | 40 ± 2.5% | 39.9333 | PASS
[7 empty] | Passed Acceptance Criteria? = | TRUE

## QS 16S Analysis
2025-11-05_095412_DEN078_RNaseP-16S.eds
Standard Curve Analysis
| Tabulated Averages | [13 empty] | Pass/Fail?
[6 empty] | Actual Results |  |  |  | Expected Results
| Sample No. | Sample ID | 1/10 Cp | Conc. (ng/μL)* | Cp SEM | Standard | Mean Cp | Mean Conc. (ng/mL)* |  | Standard | Mean Cp* | Mean Conc. (ng/mL) |  |  | PASSED
| 1 | PD005745RME2_D_Hu_01 | 20.918 | 1.3522 | 0.036986 | S1 | 13.1283 | 19397.465 |  | S1 | 12 | 20000
| 2 | PD005599RME1_D_Hu_01 | 18.9483 | 4.7465 | 0.069334 | S2 | 16.202 | 2733.6508 |  | S2 | 15 | 2500
| 3 | PD005595RME1_D_Hu_01 | 25.8723 | 0.057455 | 0.19932 | S3 | 19.6787 | 297.9635 |  | S3 | 19 | 312.5
| 4 | PD005590RME1_D_Hu_01 | 19.0653 | 4.4053 | 0.040206 | S4 | 22.935 | 37.3754 |  | S4 | 22 | 39.06
| 5 | PD005585RME1_D_Hu_01 | 21.037 | 1.2534 | 0.13434 | S5 | 26.0867 | 5.0117 |  | S5 | 25 | 4.88
| 6 | PD005594RME1_D_Hu_01 | 25.004 | 0.099941 | 0.03337 | S6 | 29.3803 | 0.61387 |  | S6 | 29 | 0.61
| 7 | PD005593RME1_D_Hu_01 | 18.201 | 7.6433 | 0.14943 | NTC | 0 | 0 |  | NTC | >35.5 | 0
| 8 | PD005592RME1_D_Hu_01 | 17.7447 | 10.2241 | 0.034439 | *Based on the Following: |  |  |  | *Based on the Following:
| 9 | PD005584RME1_D_Hu_01 | 26.0713 | 0.05061 | 0.024066 | Cp = ('Slope'*'LogConc')+'Y-Intercept'
| 10 | PD005578RME1_D_Hu_01 | 17.7707 | 10.0561 | 0.027686 | Conc (ng/mL) = 10^(('Mean Cp'-'Y-Intercept')/'Slope')
| 11 | PD005575RME1_D_Hu_01 | 19.0377 | 4.4837 | 0.03225
| 12 | PD005572RME1_D_Hu_01 | 15.9493 | 32.1143 | 0.046044
| 13 | PD005571RME1_D_Hu_01 | 19.5157 | 3.3059 | 0.044876
| 14 | PD005568RME1_D_Hu_01 | 20.3447 | 1.9488 | 0.0047687 | Standard Curve Breakdown
| 15 | PD005565RME1_D_Hu_01 | 22.1223 | 0.62746 | 0.013303 | Actual Results | [5 empty] | Full Actual Results
| 16 | PD005564RME4_D_Hu_01 | 16.2177 | 27.0648 | 0.058226 | Standard | Mean Cp | Mean Conc. (ng/mL) | Log Conc.* |  |  | Standard | Cp 1 | Cp 2 | Cp 3 | SEM | ΔCt
| 17 | PD005563RME1_D_Hu_01 | 20.5547 | 1.7046 | 0.10322 | S1 | 13.1283 | 19397.465 | 4.301 |  |  | S1 | 13.061 | 13.081 | 13.243 | 0.047049 | -
| 18 | PD005562RME1_D_Hu_01 | 19.6013 | 3.1302 | 0.087566 | S2 | 16.202 | 2733.6508 | 3.3979 |  |  | S2 | 16.265 | 16.202 | 16.139 | 0.029698 | 3.0737
| 19 | PD005561RME1_D_Hu_01 | 18.9827 | 4.6437 | 0.0051926 | S3 | 19.6787 | 297.9635 | 2.4949 |  |  | S3 | 19.777 | 19.597 | 19.662 | 0.042969 | 3.4767
| 20 | PD005558RME1_D_Hu_01 | 21.2743 | 1.0774 | 0.26647 | S4 | 22.935 | 37.3754 | 1.5917 |  |  | S4 | 22.983 | 22.979 | 22.843 | 0.037571 | 3.2563
| 21 | PD005557RME1_D_Hu_01 | 19.4707 | 3.4021 | 0.19984 | S5 | 26.0867 | 5.0117 | 0.68842 |  |  | S5 | 26.128 | 26.134 | 25.998 | 0.036226 | 3.1517
| 22 | PD005556RME1_D_Hu_01 | 20.6757 | 1.5781 | 0.10232 | S6 | 29.3803 | 0.61387 | -0.21467 |  |  | S6 | 29.436 | 29.219 | 29.486 | 0.06691 | 3.2937
| 23 | DEN078-POS_D_Hu_01 | 29.8033 | 0.0046877 | 0.14158 | *Based on Expected Concentration
| 24 | DEN078-NEG_D_Hu_01 | 29.7377 | 0.0048881 | 0.22679
| NEG Control | NF-Water | 35.291 | 0.014177 | 0.14012 | Graphical Representation
*Actual Sample Concentration, Calculated as 10x LightCycler Output
| ◊Actual Sample Concentration, Calculated as 100x LightCycler Output | [14 empty] | Plot Analysis
| Conditional Formatting Key | [14 empty] | 28.6149 | = Y-intercept
| Text | Cp too Low/Conc too High, Result may be Unreliable | [13 empty] | -3.6118 | = Slope
| Text | Cp too High/Conc too Low to Accurately Quantify
| Text | SEM too High to be reliable, Investigate Individual Cps | [13 empty] | 1.8918 | = PCR Efficiency
[16 empty] | 89.1771 | = %PCR Efficiency
[16 empty] | 0.99982 | = R Squared
Acceptance Criteria
[6 empty] | Criteria | Acceptance Range | Results | Pass/Fail?
[6 empty] | PCR Efficiency Result | 85%-115% | 89.1771 | PASS
[6 empty] | R^2 | ≥0.98 | 0.99982 | PASS
[6 empty] | NTC | Cp ≥ 35.5 | 0 | PASS
[6 empty] | Y-intercept | 28 ± 2.5% | 28.6149 | PASS
[7 empty] | Passed Acceptance Criteria? = | TRUE

## Result Export
Sample ID | Batch ID | RNase P Concentration (ng/μL) | RNase P Mean CP | RNase P CP SEM | RNase P R^2 | RNase P % PCR Efficiency | RNase P Y-Intercept | RNase P Slope | 16S Concentration (ng/μL) | 16S Mean CP | 16S CP SEM | 16S R^2 | 16S % PCR Efficiency | 16S Y-Intercept | 16S Slope | Fast Advanced MM Lot | RNase P Primer Probe Lot | Promega Lot | Human DNA Male Lot | 16S Primer Probe Lot | Microbial Community Standard Lot | NF-water Lot | Asset ID
PD005745RME2_D_Hu_01 | DEN078 | 1.1308 | 32.7407 | 0.14231 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 1.3522 | 20.918 | 0.036986 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005599RME1_D_Hu_01 | DEN078 | 2.0798 | 31.8137 | 0.053144 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 4.7465 | 18.9483 | 0.069334 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005595RME1_D_Hu_01 | DEN078 | 6.1515 | 30.164 | 0.14329 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 0.057455 | 25.8723 | 0.19932 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005590RME1_D_Hu_01 | DEN078 | 7.59 | 29.8443 | 0.062006 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 4.4053 | 19.0653 | 0.040206 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005585RME1_D_Hu_01 | DEN078 | 18.6708 | 28.475 | 0.13691 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 1.2534 | 21.037 | 0.13434 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005594RME1_D_Hu_01 | DEN078 | 4.3542 | 30.6897 | 0.056697 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 0.099941 | 25.004 | 0.03337 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005593RME1_D_Hu_01 | DEN078 | 0.87738 | 33.1267 | 0.5717 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 7.6433 | 18.201 | 0.14943 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005592RME1_D_Hu_01 | DEN078 | 6.1894 | 30.1547 | 0.040225 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 10.2241 | 17.7447 | 0.034439 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005584RME1_D_Hu_01 | DEN078 | 52.8433 | 26.8923 | 0.018893 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 0.05061 | 26.0713 | 0.024066 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005578RME1_D_Hu_01 | DEN078 | 8.5658 | 29.6603 | 0.037074 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 10.0561 | 17.7707 | 0.027686 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005575RME1_D_Hu_01 | DEN078 | 8.3949 | 29.691 | 0.022896 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 4.4837 | 19.0377 | 0.03225 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005572RME1_D_Hu_01 | DEN078 | 13.6574 | 28.9507 | 0.031164 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 32.1143 | 15.9493 | 0.046044 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005571RME1_D_Hu_01 | DEN078 | 2.0735 | 31.8183 | 0.076065 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 3.3059 | 19.5157 | 0.044876 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005568RME1_D_Hu_01 | DEN078 | 0.81618 | 33.2367 | 0.1561 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 1.9488 | 20.3447 | 0.0047687 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005565RME1_D_Hu_01 | DEN078 | 0.40857 | 34.2893 | 0.061573 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 0.62746 | 22.1223 | 0.013303 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005564RME4_D_Hu_01 | DEN078 | 3.0206 | 31.246 | 0.040108 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 27.0648 | 16.2177 | 0.058226 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005563RME1_D_Hu_01 | DEN078 | 1.6208 | 32.193 | 0.083395 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 1.7046 | 20.5547 | 0.10322 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005562RME1_D_Hu_01 | DEN078 | 1.3614 | 32.4583 | 0.05762 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 3.1302 | 19.6013 | 0.087566 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005561RME1_D_Hu_01 | DEN078 | 7.9946 | 29.7653 | 0.059506 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 4.6437 | 18.9827 | 0.0051926 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005558RME1_D_Hu_01 | DEN078 | 4.8255 | 30.5333 | 0.1269 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 1.0774 | 21.2743 | 0.26647 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005557RME1_D_Hu_01 | DEN078 | 4.4585 | 30.6537 | 0.13941 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 3.4021 | 19.4707 | 0.19984 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
PD005556RME1_D_Hu_01 | DEN078 | 1.6369 | 32.178 | 0.070225 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 1.5781 | 20.6757 | 0.10232 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
DEN078-POS_D_Hu_01 | DEN078 | 0 | 0 | 0 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 0.0046877 | 29.8033 | 0.14158 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131
DEN078-NEG_D_Hu_01 | DEN078 | 0 | 0 | 0 | 0.99951 | 92.9674 | 39.9333 | -3.5028 | 0.0048881 | 29.7377 | 0.22679 | 0.99982 | 89.1771 | 28.6149 | -3.6118 | 4444554 (Verified) | 3245729 (Verified) | 0000636473 (Verified) | 3106717 (Verified) | P250904-011-F09 (Verified) | 24027229 (Verified) | N/A | 00131

## Revision History (hidden)
Revision | Change Detail | Rationale
A | Introduction of new experiment report file | New procedure
B | Added new tabs for processing of data from the QuantStudio | Raw data files from the LightCycler and QuantStudio have different formats, meaning that the pre-existing formulas were incompatible with the QuantStudio files.
C | Combined RNase P and 16S experiment report files | New procedure